|  | Urgent Field Safety Notice Microplates E.M.®Technology for ABO blood grouping and Rh-Kell phenotyping. Potential False positive results Immediate actions required |
| --- | --- |

Our reference: FSN/2019/4

11 february 2019

Attention: Laboratory Managers, Director of Health Institutes and Vigilance Representative.

Dear valued customers

This notice has been issued to advise you of immediate actions required related to the following affected devices:

| Device Name | Reference | Lot number |
| --- | --- | --- |
| DuoLys | 79960 | see attached list |
| ABD-Lys | 79967 |  |
| Groupa 2 Lys | 79964 |  |
| Groupa A1 Lys | 79965 |  |
| Groupa DVI Lys | 79969 |  |
| PhenoLys | 79962 |  |

Problem description

We have been informed of an increase of some doubtful to false positive reactions during the grouping ABO, and phenotyping Rh-Kell tests.

The false positive images look like double population with weak intensity (? to 2+ interpretation).

Our investigations fixed the root cause of this phenomenon and determinate the concerned lots.

1% of the microplates are partially concerned due to an issue of the production process.

No damage for the true positive reactions.

Risk assessment

With double determination, a potential false positive result is very unlikely.

However in this case:

- For the ABO antigens, a discrepancy with the reverse typing, or with the compatibility test, or the ABO blood group compatibility before a transfusion at patient’s bedside, will stop the transfusion.

- For the systems Rh-Kell antigens, the risk is the immunization of patient or pregnant woman which will be detected by the antibody screening test before the next transfusion or the following pregnancy.

If only one determination is practiced for the Rh-Kell systems, the risk is a patient immunization after the first transfusion which will be detected by the antibody screening test or/and by the compatibility test before a next transfusion.

In the case of laboratories for blood donation qualification, if a blood bag with a false positive result is transfused, there is no risk for the patient to be transfused as the antigen is really absent from donor red blood cells.

Resolution

Corrective actions are implemented, and new lots are being produced. These will be progressively available.

Required actions by users

In case of doubtful to 2+ reactions, we advise you to confirm the results with another microplate or with another technique.

Please complete the enclosed Field Safety Corrective Action (FSCA) response form  	immediately and return it by fax to +33 (0)3 20 96 53 64 or by email to 	hotline@diagast.com.

Transmission of this field safety notice

Please provide a copy of this notice to all individuals within your organization, as well as to any third parties with whom you interact, who may have access to or knowledge of affected reagent. Please maintain awareness to this field safety notice as appropriate to ensure its effectiveness.

The undersigned confirm that DIAGAST has provided a copy of this FSN to the relevant competent authorities.

We regret any inconvenience that this action may have caused to you or your staff.

Sincerely yours.

Jean-Michel BASSE

Quality & Regulatory affairs Director

Attachment : FSCA response form.

List of the concerned devices lots

| Device Name | Reference | Lot number | Expiry Date |
| --- | --- | --- | --- |
| DuoLys | 79960 | 349000 | 30/11/2019 |
|  |  | 350000 | 30/11/2019 |
|  |  | 351000 | 31/12/2019 |
|  |  | 352000 | 31/12/2019 |
|  |  | 353000 | 31/12/2019 |
|  |  | 354000 | 31/01/2020 |
|  |  | 356000 | 30/04/2020 |
|  |  | 359000 | 31/05/2020 |
| ABD-Lys | 79967 | 175000 | 31/01/2020 |
|  |  | 176000 | 31/01/2020 |
|  |  | 177000 | 29/02/2020 |
|  |  | 178000 | 31/03/2020 |
|  |  | 179000 | 31/03/2020 |
|  |  | 180000 | 31/03/2020 |
|  |  | 181000 | 31/03/2020 |
|  |  | 182000 | 30/04/2020 |
|  |  | 183000 | 31/03/2020 |
|  |  | 184000 | 31/03/2020 |
|  |  | 185000 | 31/05/2020 |
| Groupa 2 Lys | 79964 | 076000 | 31/12/2019 |
|  |  | 077000 | 31/03/2020 |
|  |  | 078000 | 31/05/2020 |
|  |  | 079000 | 31/05/2020 |
| Groupa A1 Lys | 79965 | 034000 | 30/06/2019 |
| Groupa DVI Lys | 79969 | 052000 | 31/05/2019 |
| PhenoLys | 79962 | 632000 | 30/04/2020 |

|  | Field Safety Corrective Action (FSCA) Response Form Note: Please complete this form even if you do not have any remaining affected product in your inventory and return by email at hotline@diagast.com |
| --- | --- |

Our reference°: FSN /2019/4

Date: 11 february 2019

Affected devices

| Device Name | Reference | Lot number |
| --- | --- | --- |
| DuoLys | 79960 | see attached list |
| ABD-Lys | 79967 |  |
| Groupa 2 Lys | 79964 |  |
| Groupa A1 Lys | 79965 |  |
| Groupa DVI Lys | 79969 |  |
| PhenoLys | 79962 |  |

Customer information

| Customer Account Nr |  |
| --- | --- |
| Society Name |  |
| Manager Name |  |
| Adresse |  |
| Email |  |

We acknowledge the receipt of the FSCA referenced above and that the information therein has been shared with all users of the affected devices within our organization, as well as with any third parties to whom we may have transferred any affected devices.

Response date : 	Stamp/Signature :