Relevant Country Contact Name

Tel.: Relevant country telephone

Fax: Relevant country fax

Email: Relevant country email

Urgent Field Safety Notice

Devices:

| REF | DESCRIPTION |
| --- | --- |
| 1898000 | DOUBLE SWIVEL ELBOW 15M - DOUBLE FLIP TOP CAP WITH SEAL - 22M/15F |
| 3521000 | SUPERSET DOUBLE SWIVEL CATHETER MOUNT 22F - DOUBLE FLIP TOP CAP - 22M/15F, 70MM-150MM |

LOT numbers:

| 1898000 | 2180346, 2180563, 2181793, 2182007, 2190034, 2190139, 2190292, 2190976 |
| --- | --- |
| 3521000 | 2180366, 2180397, 2180571, 2180859, 2181350, 2181503, 2181825, 2182047, 2182119, 2182417, 2182852, 2183480, 2183518, 2183729, 2190008, 2190211, 2190350, 2190531, 2190735, 2190850, 2191382, 2191591 |

Manufacturer: Intersurgical Ltd

FSCA-identifier: 226766		Date:  15/07/2019

Attention: Medical Device Safety Officers (MDSO)

Distribution:  All Critical Care clinical staff, all Anaesthesia clinical staff, managers and users of the above products

Type of action: Users of the products and lot numbers listed above must follow the additional instructions described below before use.

Description of the problem: We have received complaints about cracked ports in the swivel elbow where the flip cap is inserted, and these cracks leak if they reach a significant size. The percentage of cracked products in any lot may vary between 0% and approximately 10%. The cracks were not present at the time of manufacture, but have occurred later on during storage.

Transmission of this Field Safety Notice:

This notice should be transmitted to all those who need to be aware within your organisation, or to any organisation where these potentially affected devices have been transferred.

Intersurgical apologises for any inconvenience this may cause. If you have any questions, please contact your distributor or local Intersurgical representative.

The relevant National Authorities have been advised about this Field Safety Corrective Action.

Please maintain awareness of this Field Safety Notice and resulting action for an appropriate period to ensure effectiveness of the corrective action.

Action to be taken by the user:

Visually check the swivel elbow port is not cracked prior to use, checking all around the port, and where possible the swivel elbow should be included in the pre-use leak test. Do not use if the port is cracked. Contact Intersurgical to request credit or replacement products.

Corrective Action being taken by manufacturer Intersurgical:

We have stopped production of the affected component where the problem has been identified, however in the meantime we are able to use components manufactured at a different site to minimise disruption in supply to our customers.

The undersigned confirms that this notice has been notified to the appropriate Regulatory Agency.

Ivan Seniut, Group Quality and Regulatory Affairs Director, Intersurgical

Relevant Country Contact Name

Tel.: Relevant country telephone

Fax: Relevant country fax

Email: Relevant country email

Urgent Field Safety Notice

# Response Form

Devices:

| REF | DESCRIPTION |
| --- | --- |
| 1898000 | DOUBLE SWIVEL ELBOW 15M - DOUBLE FLIP TOP CAP WITH SEAL - 22M/15F |
| 3521000 | SUPERSET DOUBLE SWIVEL CATHETER MOUNT 22F - DOUBLE FLIP TOP CAP - 22M/15F, 70MM-150MM |

LOT numbers:

| 1898000 | 2180346, 2180563, 2181793, 2182007, 2190034, 2190139, 2190292, 2190976 |
| --- | --- |
| 3521000 | 2180366, 2180397, 2180571, 2180859, 2181350, 2181503, 2181825, 2182047, 2182119, 2182417, 2182852, 2183480, 2183518, 2183729, 2190008, 2190211, 2190350, 2190531, 2190735, 2190850, 2191382, 2191591 |

Manufacturer: Intersurgical Ltd

FSCA-identifier: 226766		Date:  15/07/2019

Hospital Name: 	 ________________________________________

Hospital Address:	 ________________________________________

Please tick one of the following options:

We do not have any remaining stock of the affected products.

We have read the additional instructions contained within this FSN and are happy to continue using the products we have in our stock.

We have quarantined our remaining stock of the following affected products and wish to return them for credit (Please complete the section below, and send it back to the contact above).

I confirm that I have quarantined the following products and lot numbers.

| REF | LOT | Quantity of products per LOT number |
| --- | --- | --- |
| [add more rows as required] |  |  |

Form Completed and Returned by:

Name: …………………………………………..		Phone No: ………………………..

Position: ………………………………………. 		E-mail: ……………………………

Date: ……………………………………………