| To: | Name/name: Abtlg./Dept. Tel./phone: Email: Datum/date : | Tom Peters Quality Management +49 8153 888-155 mdro@dornier.com Wessling, 05.04.2019 |
| --- | --- | --- |

Urgent Field Safety Notice (FSN):

Shelf Life Issue of Dornier Standard Diode and Nd:YAG Lightguides

Dear Valued Customer,

The purpose of this letter is to advise you that Dornier MedTech, the manufacturer of the Dornier Diode and Nd:YAG Laser Fibers, is voluntarily issuing a Medical Device Recall regarding the laser fibers listed in the table below.  Distribution records indicate that you have received one or more of these products that are the subject of this action.

During a routine re-evaluation of the packaging design for the diode fibers conducted to confirm a 5-year shelf life, test results indicated that the current package design, paper/poly pouch, of the products with the article numbers as listed below, showed pin hole package failures.  The samples tested were units retrieved from the field during the aforesaid routine re-evaluation.  Dornier has not received any customer complaints regarding this issue.  Although further testing is underway, out of an abundance of caution, Dornier has stopped shipping all fibers that have this package design and has initiated this action.  Accordingly, Dornier MedTech is issuing this Field Safety Notice as a precautionary measure to prevent the use of product where the sterility barrier may have been compromised.

The Dornier Diode Laser Fibers are intended to be used as an accessory for Dornier Medilas D Family of Lasers, as well as the Nd:YAG Family of General Surgical Lasers.

The table below lists the affected products:

| Article No | Name |  | Article No | Name |
| --- | --- | --- | --- | --- |
| K1008084 | LL-E D01-6100-BF-0 10ST BAREFIBER |  | K2012393 | LL-E D01-6080-BF-0 10ST BAREFIBER |
| K1009920 | LL-E D01-6180-D-0 10ST DUESE 1.8MM |  | K2012867 | LL-E S00-4079-BF-0 10ST BAREFIBER |
| K1009922 | LL-E D01-6100-B-0 10ST BAREFIBER SPUEL |  | K2012870 | LL-E S00-6109-BF-0 10ST BAREFIBER |
| K1009924 | LL-E D01-6210-D-0 10ST DUESE 2.1MM |  | K2013073 | LL-E S02-6080-BF-1 10ST BAREFIBER |
| K2010292 | LL-E D01-4070-BF-1 10ST BAREFIBER HCL |  | K1001291 | LL-EINM ND E-6210-D 10ST DUESE 2.1MM |
| K2010710 | LL-E D01-4070-BF-0 10ST BAREFIBER HCP |  | K1001292 | LL-EINM ND E-6180-D 10ST DUESE 1.8MM |
| K2011577 | LL-E S01-4070-BF-0 10ST BAREFIBER HCP |  | K1001295 | LL-EINM ND E-6100-B 10ST BAREFIB SPUEL |
| K2011580 | LL-E S01-6100-BF-0 10ST BAREFIBER |  | K1001296 | LL-EINM ND E-4070-B 10ST BAREFIB SPUEL |
| K2011594 | LL-E S01-6180-D-0 10ST DUESE 1.8MM |  | K1010498 | LL-EINM ND E-4070-BF 10ST BAREFIBER |
| K2011596 | LL-E S01-6210-D-0 10ST DUESE 2.1MM |  | K1001293 | LL-EINM ND E-6220-G 10ST GEWINDE 2.2MM |
| K2011822 | LL-E D00-4079-BF-0 10ST BAREFIBER |  | K1001294 | LL-EINM ND E-6180-G 10ST GEWINDE 1.8MM |
| K2011824 | LL-E D00-6109-BF-0 10ST BAREFIBER |  | K1010500 | LL-EINM ND E-6100-BF 10ST BAREFIBER |
| K2012146 | LL-E S02-6100-BF-0 10ST BAREFIBER |  |  |  |

Pin hole package failures may lead to fibers being unsterile. Given a contaminated laser fiber was used during a clinical procedure, the health consequences mainly consist of microbial contamination of the surgical site that could lead to surgical site infection. Such events could be result in patient pain and the potential for increased length of hospital stay. Such infection could be treated (e.g., antibiotics).  Long-term manifestations, although rare, could include osteomyelitis, soft tissue abscess, and sepsis with worst case being septic shock if not treated. Dornier MedTech is not aware of any clinical incidents related to this packaging issue and no customer complaints have been reported.

What does the user have to do?

- Identify if you have above listed products in stock (article no., lot, quantity).

- If you have above listed products in stock, please immediately quarantine all inventory of Dornier Diode and Nd:YAG laser fibers and do not use these products until further notice to avoid any health risks to the patient.

- We are currently analyzing the situation and will inform you what to do with your quarantined stock of lightguides once that determination has been made.

- If you seek a replacement for your fibers, please check the box on the feedback sheet and send us your affected product. Dornier MedTech will then contact you regarding suitable replacements.

- We further on request you to please fill in the attached form immediately and send it back within 7 calendar days by FAX or e-Mail, to below mentioned contact.

Thank you for your support.  Please contact your local sales contact if you have any questions regarding this FSN, any of our products, or would like assistance with the action:

Dornier MedTech Europe GmbH

Orest Buchma

Argelsrieder Feld 7, 82334 Wessling

Tel.: +49 8153 888-548

Mail: obuchma@dornier.com

We regret any inconvenience that this action may cause, but we appreciate your understanding as we take steps to ensure patient and customer satisfaction.

Kind regards

Dornier MedTech GmbH

Tom Peters

Medical Device Reporting Officer